Design WS1050 Reliability Digital I/O

36 DUT Board Layout:

Test Setup Block Diagram:

Hardware:

- Keithley 2000 (DMM)

- Keithley 2400 (Source meter for VSA test)

- Agilent Power supply E3631 (Power source for +3.3V, VDD)

- Chassis PXI 1033

- FPGA card Rio 7813R, 160 digital I/O channels

- SHC68-68-RDIO Cables (4)

- Several banana plugs

Software:

- FPGA Revision :

- WS_VCAP Driver 029 (36 DUT)

- WS_VCAP Driver 024 (16 DUT)

- Stress Test Program

- WS1050 36 DUT Reliability HD_Cycling Ver1.3 (36 DUT)

- WS1050 16 DUT Reliability HD_Cycling Ver1.8 (16 DUT)

Schematic:

- Schematic WS1050 36 DUT Reliability REV E.

- Schematic WS1050 16 DUT Reliability REV G.

Test Sequences:

- Turn ON +3.3V, this will turn on AD7747 and voltage regulator LM317 to generate 1.8V for VIO

- Turn ON VDD (HD = 3.3V; Cycle = 3.7V)

- Set CP OFF

- Measure IDD standby

- Set CP Voltage to stress

- Set DVA (ON/OFF)

- Set ECAL (ON/OFF)

- Set VPI (ON/OFF) Note: If VPI OFF, skip steps 31 to 41.

- Set VSA (ON/OFF) Note: If VSA OFF, skip steps 42 to 54.

- Set HS (ON/OFF) Note: If HS check OFF, skip step 55.

- Start HD/Cycle (Non linear RP).

- Turn All Drivers ON

- Measure IDD All Drivers ON

- Turn All Drivers OFF

- Measure IDD All Drivers OFF

- Measure C1_Min

- Measure C1_F1

- Measure C1_F2

- Measure C1_S1 to C1_S14F12

- Measure C1_Min2

- Measure C2_Min

- Measure C2_F1

- Measure C2_F2

- Measure C2_S1 to C2_S14F12

- Measure C2_Min2

- Measure C3_Min

- Measure C3_F1

- Measure C3_F2

- Measure C3_S1 to C3_S14F12

- Measure C3_Min2

- Set DVA OFF, ECAL OFF (this is for VPI)

- Set CP Voltage for VPI

- Measure C1_Min

- Measure C1_Max

- Measure C1_Min2

- Measure C2_Min

- Measure C2_Max

- Measure C2_Min2

- Measure C3_Min

- Measure C3_Max

- Measure C3_Min2

- Set Ke2400 to V1 = -10V

- Measure VSA C1_V1

- Measure VSA C2_V1

- Measure VSA C3_V1

- Set Ke2400 to V2 = 0V

- Measure VSA C1_V2

- Measure VSA C2_V2

- Measure VSA C3_V2

- Set Ke2400 to V3 = 10V

- Measure VSA C1_V3

- Measure VSA C2_V3

- Measure VSA C3_V3

- Set Ke2400 to 0V.

- Run HS check on Register 1 (bank C1)

- Save Data

- Loop back to step 11

Note:

Test Sequence:  Test program can measure read point by bank order, C1-C2-C3; C2-C1-C3 or C3-C2-C1

Stress Sequence:  Test program can stress parts by bank order by setting commands:

070201|XX|XX|XX

070201|XX|XX|XX

070201|XX|XX|XX

XX can be 3B all drivers ON

XX can be 00 all drivers OFF

Cycling speed is12KHz with signal CLK is 5MHz.

With 33% Duty Cycle:

With 25% Duty Cycle

- Main GUI for calibration and set up test:

- GUI to monitor the test setup (CP voltage, DVA ON/OFF, ECAL ON/OFF) before stress, monitor cap values for each drivers, and time completed stress

- GUI to monitor VSA test, HS check, EFUSE read back  and % Delta C change

- GUI to monitor CMIN changes during the stress